Table S1. Comparison of the results obtained with three molecular-diagnostics SARS-CoV-2 assays

| Specimen number | Results for clinical specimens (cycle threshold value) |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  | Reference RT-PCR method | BioFire RP2.1 | Cobas Liat System | Cepheid Xpert | Lineage |  |
| 1 | Positive (25.0) | Positive | Positive (23.7) | Positive (26.8) | non B.1.1.7 |  |
| 2 | Positive (29.8) | Positive | Positive (28.2) | Positive (31.2) | non B.1.1.7 |  |
| 3 | Positive (34.2) | Positive | Positive (30.6) | Positive (32.3) | non B.1.1.7 |  |
| 4 | Positive (25.6) | Positive | Positive (24.5) | Positive (27.8) | non B.1.1.7 |  |
| 5 | Positive (27.2) | Positive | Positive (28.7) | Positive (33.0) | non B.1.1.7 |  |
| 6 | Positive (28.5) | Positive | Positive (28.2) | Positive (31.8) | non B.1.1.7 |  |
| 7 | Positive (32.4) | Negative | Positive (35.3) | Positive (36.8) | non B.1.1.7 |  |
| 8 | Positive (30.9) | Positive | Positive (30.9) | Positive (31.4) | non B.1.1.7 |  |
| 9 | Positive (30.3) | Positive | Positive (30.1) | Positive (33.7) | non B.1.1.7 |  |
| 10 | Positive (30.4) | Positive | Positive (27.7) | Positive (31.3) | non B.1.1.7 |  |
| 11 | Positive (26.4) | Positive | Positive (26.9) | Positive (30.7) | non B.1.1.7 |  |
| 12 | Positive (28.6) | Positive | Positive (27.7) | Positive (31.4) | non B.1.1.7 |  |
| 13 | Positive (18.3) | Positive | Positive (16.9) | Positive (19.3) | non B.1.1.7 |  |
| 14 | Positive (26.3) | Positive | Positive (26.8) | Positive (30.0) | non B.1.1.7 |  |
| 15 | Positive (27.2) | Positive | Positive (27.1) | Positive (31.1) | non B.1.1.7 |  |
| 16 | Positive (31.3) | Positive | Positive (28.2) | Positive (34.5) | non B.1.1.7 |  |
| 17 | Positive (31.8) | Positive | Positive (29.5) | Positive (35.9) | non B.1.1.7 |  |
| 18 | Positive (30.7) | Positive | Positive (32.3) | Positive (36.5) | non B.1.1.7 |  |
| 19 | Positive (29.7) | Positive | Positive (27.1) | Positive (33.1) | non B.1.1.7 |  |
| 20 | Positive (26.9) | Positive | Positive (25.6) | Positive (28.9) | non B.1.1.7 |  |
| 21 | Positive (29.9) | Positive | Positive (28.8) | Positive (32.8) | non B.1.1.7 |  |
| 22 | Positive (31.5) | Positive | Positive (31.5) | Positive (37.4) | non B.1.1.7 |  |
| 23 | Positive (31.1) | Positive | Positive (29.7) | Positive (35.8) | non B.1.1.7 |  |
| 16 | Positive (31.3) | Positive | Positive (28.2) | Positive (34.5) | non B.1.1.7 |  |
| 17 | Positive (31.8) | Positive | Positive (29.5) | Positive (35.9) | non B.1.1.7 |  |
| 18 | Positive (30.7) | Positive | Positive (32.3) | Positive (36.5) | non B.1.1.7 |  |
| 19 | Positive (29.7) | Positive | Positive (27.1) | Positive (33.1) | non B.1.1.7 |  |
| 20 | Positive (26.9) | Positive | Positive (25.6) | Positive (28.9) | non B.1.1.7 |  |
| 21 | Positive (29.9) | Positive | Positive (28.8) | Positive (32.8) | non B.1.1.7 |  |
| 22 | Positive (31.5) | Positive | Positive (31.5) | Positive (37.4) | non B.1.1.7 |  |
| 23 | Positive (31.1) | Positive | Positive (29.7) | Positive (35.8) | non B.1.1.7 |  |
| 24 | Positive (31.9) | Positive | Positive (31.4) | Positive (34.3) | non B.1.1.7 |  |
| 25 | Positive (32.9) | Positive | Positive (31.2) | Positive (37.4) | non B.1.1.7 |  |
| 26 | Positive (23.7) | Positive | Positive (23.2) | Positive (27.4) | non B.1.1.7 |  |
| 27 | Positive (27.5) | Positive | Positive (27.9) | Positive (31.3) | non B.1.1.7 |  |
| 28 | Positive (26.2) | Positive | Positive (25.3) | Positive (30.2) | non B.1.1.7 |  |
| 29 | Positive (30.1) | Positive | Positive (28.3) | Positive (31.4) | non B.1.1.7 |  |
| 30 | Positive (28.4) | Positive | Positive (25.7) | Positive (31.0) | non B.1.1.7 |  |
| 31 | Positive (30.9) | Positive | Positive (31.9) | Positive (34.5) | non B.1.1.7 |  |
| 32 | Positive (31.3) | Positive | Positive (29.4) | Positive (32.9) | non B.1.1.7 |  |
| 33 | Positive (32.7) | Positive | Positive (32.4) | Positive (37.2) | non B.1.1.7 |  |
| 34 | Positive (32.1) | Positive | Positive (30.4) | Positive (36.6) | non B.1.1.7 |  |
| 35 | Positive (35.4) | Positive | Positive (30.6) | Positive (35.1) | non B.1.1.7 |  |
| 36 | Positive (31.7) | Positive | Positive (30.3) | Positive (35.3) | non B.1.1.7 |  |
| 37 | Positive (31.4) | Positive | Positive (30.1) | Positive (36.7) | non B.1.1.7 |  |
| 38 | Positive (31.5) | Positive | Positive (29.5) | Positive (33.8) | non B.1.1.7 |  |
| 39 | Positive (25.5) | Positive | Positive (24.7) | Positive (29.3) | non B.1.1.7 |  |
| 40 | Positive (30.5) | Positive | Positive (30.1) | Positive (34.2) | non B.1.1.7 |  |
| 41 | Positive (23.7) | Positive | Positive (22.1) | Positive (25.9) | non B.1.1.7 |  |
| 42 | Positive (19.0) | Positive | Positive (18.6) | Positive (20.1) | non B.1.1.7 |  |
| 43 | Positive (19.3) | Positive | Positive (19.2) | Positive (22.7) | non B.1.1.7 |  |
| 44 | Positive (21.3) | Positive | Positive (23.2) | Positive (25.8) | non B.1.1.7 |  |
| 45 | Positive (26.5) | Positive | Positive (26.1) | Positive (28.4) | non B.1.1.7 |  |
| 46 | Positive (20.5) | Positive | Positive (22.1) | Positive (26.4) | non B.1.1.7 |  |
| 47 | Positive (21.3) | Positive | Positive (22.3) | Positive (25.1) | non B.1.1.7 |  |
| 48 | Positive (21.1) | Positive | Positive (20.1) | Positive (23.3) | non B.1.1.7 |  |
| 49 | Positive (24.5) | Positive | Positive (23.2) | Positive (28.6) | non B.1.1.7 |  |
| 50 | Positive (21.4) | Positive | Positive (21.2) | Positive (23.6) | non B.1.1.7 |  |
| 51 | Positive (20.8) | Positive | Positive (24.1) | Positive (24.7) | B.1.1.7 |  |
| 52 | Positive (23.8) | Positive | Positive (27.2) | Positive (27.8) | B.1.1.7 |  |
| 51 | Positive (26.2) | Positive | Positive (30.5) | Positive (30.6) | B.1.1.7 |  |
| 52 | Positive (26.8) | Positive | Positive (30.5) | Positive (31.8) | B.1.1.7 |  |
| 53 | Positive (28.1) | Positive | Positive (31.5) | Positive (32.4) | B.1.1.7 |  |
| 54 | Positive (17.7) | Positive | Positive (19.7) | Positive (21.4) | B.1.1.7 |  |
| 55 | Positive (23.8) | Positive | Positive (25.8) | Positive (27.3) | B.1.1.7 |  |
| 56 | Positive (26.3) | Positive | Positive (28.2) | Positive (30.5) | B.1.1.7 |  |
| 57 | Positive (28.7) | Positive | Positive (30.1) | Positive (33.7) | B.1.1.7 |  |
| 58 | Positive (28.9) | Positive | Positive (30.1) | Positive (32.0) | B.1.1.7 |  |
| 59 | Positive (21.5) | Positive | Positive (24.3) | Positive (25.6) | B.1.1.7 |  |
| 60 | Positive (25.6) | Positive | Positive (30.2) | Positive (30.0) | B.1.1.7 |  |
| 61 | Positive (28.5) | Positive | Positive (32.3) | Positive (33.8) | B.1.1.7 |  |
| 62 | Positive (28.1) | Positive | Positive (31.5) | Positive (33.1) | B.1.1.7 |  |
| 63 | Positive (24.0) | Positive | Positive (25.9) | Positive (25.8) | B.1.1.7 |  |
| 64 | Positive (26.3) | Positive | Positive (28.3) | Positive (28.4) | B.1.1.7 |  |
| 65 | Positive (28.7) | Positive | Positive (30.4) | Positive (30.6) | B.1.1.7 |  |
| 66 | Positive (30.8) | Positive | Positive (32.3) | Positive (32.4) | B.1.1.7 |  |
| 67 | Positive (19.5) | Positive | Positive (21.4) | Positive (21.5) | B.1.1.7 |  |
| 68 | Positive (23.0) | Positive | Positive (24.9) | Positive (25.2) | B.1.1.7 |  |
| 69 | Positive (26.1) | Positive | Positive (27.8) | Positive (28.2) | B.1.1.7 |  |
| 70 | Positive (28.3) | Positive | Positive (30.3) | Positive (30.9) | B.1.1.7 |  |
| 71 | Positive (29.9) | Positive | Positive (32.1) | Positive (32.1) | B.1.1.7 |  |
| 72 | Positive (20.5) | Positive | Positive (23.8) | Positive (24.5) | B.1.1.7 |  |
| 73 | Positive (24.5) | Positive | Positive (28.2) | Positive (28.7) | B.1.1.7 |  |
| 74 | Positive (26.8) | Positive | Positive (30.0) | Positive (30.6) | B.1.1.7 |  |
| 75 | Positive (30.6) | Positive | Positive (33.3) | Positive (33.5) | B.1.1.7 |  |
| 76 | Positive (19.3) | Positive | Positive (22.3) | Positive (23.1) | B.1.1.7 |  |
| 77 | Positive (23.8) | Positive | Positive (27.0) | Positive (27.8) | B.1.1.7 |  |
| 78 | Positive (26.5) | Positive | Positive (30.6) | Positive (30.7) | B.1.1.7 |  |
| 79 | Positive (29.3) | Positive | Positive (31.1) | Positive (32.8) | B.1.1.7 |  |
| 80 | Positive (19.7) | Positive | Positive (22.6) | Positive (23.2) | B.1.1.7 |  |
| 81 | Positive (23.9) | Positive | Positive (27.1) | Positive (27.4) | B.1.1.7 |  |
| 82 | Positive (27.1) | Positive | Positive (30.6) | Positive (31.0) | B.1.1.7 |  |
| 83 | Positive (28.7) | Positive | Positive (31.5) | Positive (31.6) | B.1.1.7 |  |
| 84 | Positive (19.8) | Positive | Positive (23.0) | Positive (23.5) | B.1.1.7 |  |
| 85 | Positive (22.5) | Positive | Positive (25.6) | Positive (26.2) | B.1.1.7 |  |
| 86 | Positive (27.5) | Positive | Positive (30.3) | Positive (31.2) | B.1.1.7 |  |
| 87 | Positive (29.4) | Positive | Positive (32.8) | Positive (34.0) | B.1.1.7 |  |
| 88 | Positive (17.0) | Positive | Positive (20.3) | Positive (20.9) | B.1.1.7 |  |
| 89 | Positive (19.4) | Positive | Positive (22.3) | Positive (23.0) | B.1.1.7 |  |
| 90 | Positive (23.9) | Positive | Positive (26.9) | Positive (27.9) | B.1.1.7 |  |
| 91 | Positive (26.7) | Positive | Positive (28.9) | Positive (31.0) | B.1.1.7 |  |
| 92 | Positive (28.0) | Positive | Positive (30.8) | Positive (32.8) | B.1.1.7 |  |
| 93 | Positive (20.2) | Positive | Positive (23.1) | Positive (23.8) | B.1.1.7 |  |
| 94 | Positive (19.1) | Positive | Positive (22.2) | Positive (22.7) | B.1.1.7 |  |
| 95 | Positive (21.7) | Positive | Positive (24.6) | Positive (25.3) | B.1.1.7 |  |
| 96 | Positive (24.7) | Positive | Positive (27.7) | Positive (28.2) | B.1.1.7 |  |
| 97 | Positive (27.4) | Positive | Positive (30.7) | Positive (32.0) | B.1.1.7 |  |
| 98 | Positive (19.6) | Positive | Positive (19.6) | Positive (19.9) | B.1.1.7 |  |
| 99 | Positive (19.7) | Positive | Positive (20.3) | Positive (20.4) | B.1.1.7 |  |
| 100 | Positive (20.1) | Positive | Positive (20.4) | Positive (20.7) | B.1.1.7 |  |
| 101 | Positive (20.3) | Positive | Positive (20.9) | Positive (21.0) | B.1.1.7 |  |
| 102 | Positive (20.6) | Positive | Positive (20.7) | Positive (20.9) | B.1.1.7 |  |
| 103 | Positive (20.5) | Positive | Positive (20.7) | Positive (20.9) | B.1.1.7 |  |
| 104 | Positive (20.7) | Positive | Positive (20.9) | Positive (21.8) | B.1.1.7 |  |
| 105 | Positive (20.5) | Positive | Positive (20.7) | Positive (20.9) | B.1.1.7 |  |
| 106 | Positive (20.7) | Positive | Positive (20.8) | Positive (20.9) | B.1.1.7 |  |
| 107 | Positive (21.0) | Positive | Positive (21.3) | Positive (21.6) | B.1.1.7 |  |
| 108 | Positive (21.1) | Positive | Positive (21.7) | Positive (21.8) | B.1.1.7 |  |
| 109 | Positive (21.1) | Positive | Positive (21.3) | Positive (21.9) | B.1.1.7 |  |
| 110 | Positive (21.6) | Positive | Positive (22.3) | Positive (22.4) | B.1.1.7 |  |
| 111 | Positive (22.4) | Positive | Positive (22.4) | Positive (22.8) | B.1.1.7 |  |
| 112 | Positive (22.4) | Positive | Positive (22.9) | Positive (23.8) | B.1.1.7 |  |
| 113 | Positive (23.0) | Positive | Positive (23.7) | Positive (24.1) | B.1.1.7 |  |
| 114 | Positive (24.9) | Positive | Positive (25.0) | Positive (25.1) | B.1.1.7 |  |
| 115 | Positive (25.7) | Positive | Positive (26.0) | Positive (26.1) | B.1.1.7 |  |
| 116 | Positive (26.0) | Positive | Positive (26.6) | Positive (27.1) | B.1.1.7 |  |
| 117 | Positive (26.8) | Positive | Positive (26.9) | Positive (27.5) | B.1.1.7 |  |
| 118 | Positive (21.5) | Positive | Positive (22.4) | Positive (22.6) | B.1.1.7 |  |
| 119 | Positive (22.5) | Positive | Positive (26.4) | Positive (24.8) | B.1.1.7 |  |
| 120 | Positive (28.1) | Positive | Positive (30.1) | Positive (31.5) | B.1.1.7 |  |
| 121 | Positive (27.1) | Positive | Positive (32.1) | Positive (29.9) | B.1.1.7 |  |
| 122 | Positive (29.0) | Positive | Positive (33.3) | Positive (32.3) | B.1.1.7 |  |
| 123 | Positive (30.3) | Positive | Positive (33.3) | Positive (33.7) | B.1.1.7 |  |
| 124 | Positive (16.7) | Positive | Positive (20.9) | Positive (19.5) | B.1.1.7 |  |
| 125 | Positive (26.9) | Positive# | Positive (26.5) | Positive (27.4) | B.1.1.7 |  |
| 126-325 | Negative | Negative | Negative | Negative | N/A |  |

# Only M gene detected but failed to detect S gene, with results interpreted as positive according the package insert instructions.