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Suppl. Table 1: The sensitivity and specificity of clinical performance using VIDAS® SARS‑CoV‑2 IgG and IgM assay on a VIDAS® instrument

	No. of days after
PCR positive result
	Number of samples
	Number of VIDAS®
positive results
	Positive percent
agreement
	95%
Confidence interval

	Sensitivity

	Sensitivity IgG 
≥ 15 days
	40
	39
	97.5%
	88.4-99.8

	Sensitivity IgM 
≥ 15 days
	40
	38
	95.0%
	87.5-99.1

	Specificity

	Specificity IgG 
≥ 15 days
	40
	40
	100.0%
	98.9-100

	Specificity IgM 
≥ 15 days
	40
	39
	95.5%
	97.0-99.6














Suppl. Table 2: Positive and negative predictive values to the disease in the population. 

	Number of Days Post
Symptoms Onset
	PPV
	NPV

	IgG, ≥ 7 days 
(95% CI)
	93.0%
(65.6-98.7)
	99.4%
(98.9-99.8)

	IgM, ≥ 7 days
(95% CI)
	92.5%
(65.2-98.9)
	99.5%
(99.0-99.7)



*PPV denotes a Positive Predictive Value, Negative Predictive Value, Confidence interval. The calculation was done with the assumption of 5% prevalence. 
















