RESULTS

Additional information for MTT Cytotoxicity

Table S1. MTT Cytotoxicity results

| Material | Percent Viability [%] | Standard Deviation [%] | System Suitability |
| --- | --- | --- | --- |
| 4SEAL® |  |  |  |
| Positive control | 2.18 | 0.36 | Valid |
| Negative control | 118.72 | 3.30 | Valid |
| Dilution 1x | 99.55 | 4.05 | no cytotoxic potential |
| Dilution 2x | 101.94 | 3.31 | no cytotoxic potential |
| Dilution 3x | 101.25 | 1.86 | no cytotoxic potential |
| Dilution 4x | 96.13 | 7.42 | no cytotoxic potential |
| Starsil® |  |  |  |
| Positive control | 2.42 | 1.58 | Valid |
| Negative control | 96.03 | 3.22 | Valid |
| Dilution 1x | 103.05 | 3.13 | no cytotoxic potential |
| Dilution 2x | 119.48 | 8.35 | no cytotoxic potential |
| Dilution 3x | 114.17 | 5.10 | no cytotoxic potential |
| Dilution 4x | 109.67 | 2.63 | no cytotoxic potential |
| PerClot® |  |  |  |
| Positive control | 0.81 | 0.06 | Valid |
| Negative control | 91.69 | 4.80 | Valid |
| Dilution 1x | 68.47 | 1.76 | cytotoxic potential |
| Dilution 2x | 82.81 | 1.45 | no cytotoxic potential |
| Dilution 3x | 86.53 | 2.39 | no cytotoxic potential |
| Dilution 4x | 88.64 | 2.33 | no cytotoxic potential |
| ARISTA™ |  |  |  |
| Positive control | 0.81 | 0.06 | Valid |
| Negative control | 91.69 | 4.80 | Valid |
| Dilution 1x | 67.29 | 0.78 | cytotoxic potential |
| Dilution 2x | 89.07 | 2.65 | no cytotoxic potential |
| Dilution 3x | 98.66 | 2.66 | no cytotoxic potential |
| Dilution 4x | 101.99 | 4.39 | no cytotoxic potential |
| SuperClot® |  |  |  |
| Positive control | 0.81 | 0.06 | Valid |
| Negative control | 91.69 | 4.80 | Valid |
| Dilution 1x | 56.06 | 1.36 | cytotoxic potential |
| Dilution 2x | 84.17 | 2.53 | no cytotoxic potential |
| Dilution 3x | 85.41 | 2.28 | no cytotoxic potential |
| Dilution 4x | 87.89 | 2.97 | no cytotoxic potential |
| 4DryField® |  |  |  |
| Positive control | 2.42 | 1.58 | Valid |
| Negative control | 96.03 | 3.22 | Valid |
| Dilution 1x | 72.60 | 6.74 | no cytotoxic potential |
| Dilution 2x | 95.52 | 0.79 | no cytotoxic potential |
| Dilution 3x | 95.93 | 0.98 | no cytotoxic potential |
| Dilution 4x | 94.54 | 1.30 | no cytotoxic potential |
| Quality check of assay | Result |  | System Suitability |
| Absolute value of optical density (OD570) | 0.34 |  | Valid |
| Systematic cell seeding errors | 1.36 % |  | Valid |

Additional information for Water Absorption Capacity (WAC)

Table S2. Results of water absorption capacity (WAC) for 0.5 g of hemostatic powders

| Specimen | Water Absorption Capacity (WAC) [mL] |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | 4SEAL® | Starsil® | PerClot® | ARISTA™ | SuperClot® | 4DryField® |
| 1 | 22.206 | 20.137 | 19.984 | 7.550 | 6.227 | 17.488 |
| 2 | 22.342 | 17.535 | 19.846 | 8.789 | 6.381 | 18.517 |
| 3 | 22.819 | 17.831 | 19.787 | 9.776 | 6.782 | 18.019 |
| Mean | 22.456 | 18.501 | 19.872 | 8.705 | 6.463 | 18.008 |
| Standard Deviation | 0.322 | 1.425 | 0.101 | 1.115 | 0.287 | 0.515 |

Table S3. Result of one-way ANOVA with Bonferroni's multiple comparisons test for Water Absorption Capacity (WAC)

| Bonferroni's multiple comparisons test | Mean Diff. | Summary | Adjusted P Value |
| --- | --- | --- | --- |
| 4Seal® vs. Starsil® | 3.955 | ** | 0.0030 |
| 4Seal® vs. PerClot® | 2.583 | ns | 0.0591 |
| 4Seal® vs. ARISTA™ | 13.75 | **** | <0.0001 |
| 4Seal® vs. SuperClot® | 15.99 | **** | <0.0001 |
| 4Seal® vs. 4DryField® | 4.448 | ** | 0.0012 |
| Starsil® vs. PerClot® | -1.371 | ns | >0.9999 |
| Starsil® vs. ARISTA™ | 9.796 | **** | <0.0001 |
| Starsil® vs. SuperClot® | 12.04 | **** | <0.0001 |
| Starsil® vs. 4DryField® | 0.4930 | ns | >0.9999 |
| PerClot® vs. ARISTA™ | 11.17 | **** | <0.0001 |
| PerClot® vs. SuperClot® | 13.41 | **** | <0.0001 |
| PerClot® vs. 4DryField® | 1.864 | ns | 0.3412 |
| ARISTA™ vs. SuperClot® | 2.242 | ns | 0.1346 |
| ARISTA™ vs. 4DryField® | -9.303 | **** | <0.0001 |
| SuperClot® vs. 4DryField® | -11.54 | **** | <0.0001 |

| Legend: |  |
| --- | --- |
| Symbol: | p-value: |
| ns | 0.1234 |
| * | 0.0332 |
| ** | 0.0021 |
| *** | 0.0002 |
| **** | <0.0001 |

Figure S1. Result of one-way ANOVA with Bonferroni's multiple comparisons test for Water Absorption Capacity (WAC).

Additional information for Water Absorption Rate (WAR)

Table S4. Results of water absorption rate (WAR) for 0.1 g of hemostatic powders

| Specimen ▼ | Water absorption rate (WAR) [mL] |  |  |
| --- | --- | --- | --- |
| Time [s] ► | 20 | 40 | 60 |
| 4Seal® |  |  |  |
| 1 | 1.5 | 2.2 | 2.5 |
| 2 | 2.2 | 2.5 | 2.6 |
| 3 | 2.2 | 2.5 | 2.6 |
| Mean | 1.97 | 2.40 | 2.57 |
| Standard Deviation | 0.40 | 0.17 | 0.06 |
| Starsil® |  |  |  |
| 1 | 1.8 | 2.1 | 2.2 |
| 2 | 2.1 | 2.6 | 2.7 |
| 3 | 1.9 | 2.3 | 2.4 |
| Mean | 1.93 | 2.33 | 2.43 |
| Standard Deviation | 0.15 | 0.25 | 0.25 |
| PerClot® |  |  |  |
| 1 | 0.5 | 0.8 | 1.0 |
| 2 | 0.5 | 0.7 | 0.9 |
| 3 | 0.5 | 0.9 | 1.0 |
| Mean | 0.50 | 0.80 | 0.97 |
| Standard Deviation | 0.00 | 0.10 | 0.06 |
| ARISTA™ |  |  |  |
| 1 | 0.6 | 0.9 | 1.0 |
| 2 | 0.6 | 0.9 | 1.0 |
| 3 | 0.5 | 0.7 | 0.8 |
| Mean | 0.57 | 0.83 | 0.93 |
| Standard Deviation | 0.06 | 0.12 | 0.12 |
| SuperClot® |  |  |  |
| 1 | 0.5 | 0.6 | 0.7 |
| 2 | 0.3 | 0.5 | 0.5 |
| 3 | 0.3 | 0.4 | 0.5 |
| Mean | 0.37 | 0.50 | 0.57 |
| Standard Deviation | 0.12 | 0.10 | 0.12 |
| 4DryField® |  |  |  |
| 1 | 1.0 | 1.2 | 1.3 |
| 2 | 0.9 | 1.0 | 1.0 |
| 3 | 1.0 | 1.1 | 1.2 |
| Mean | 0.97 | 1.10 | 1.17 |
| Standard Deviation | 0.06 | 0.10 | 0.15 |

Table S5. Result of two-way ANOVA with Bonferroni's multiple comparisons test for Water Absorption Rate (WAR)

| Bonferroni's multiple comparisons test | Mean Diff. | Summary | Adjusted P Value |
| --- | --- | --- | --- |
| 20 s |  |  |  |
| 4Seal® vs. Starsil® | 0.03333 | ns | >0.9999 |
| 4Seal® vs. PerClot® | 1.467 | **** | <0.0001 |
| 4Seal® vs. ARISTA™ | 1.400 | **** | <0.0001 |
| 4Seal® vs. SuperClot® | 1.600 | **** | <0.0001 |
| 4Seal® vs. 4DryField® | 1.000 | **** | <0.0001 |
| Starsil® vs. PerClot® | 1.433 | **** | <0.0001 |
| Starsil® vs. ARISTA™ | 1.367 | **** | <0.0001 |
| Starsil® vs. SuperClot® | 1.567 | **** | <0.0001 |
| Starsil® vs. 4DryField® | 0.9667 | **** | <0.0001 |
| PerClot® vs. ARISTA™ | -0.06667 | ns | >0.9999 |
| PerClot® vs. SuperClot® | 0.1333 | ns | >0.9999 |
| PerClot® vs. 4DryField® | -0.4667 | * | 0.0159 |
| ARISTA™ vs. SuperClot® | 0.2000 | ns | >0.9999 |
| ARISTA™ vs. 4DryField® | 0.4000 | ns | 0.0636 |
| SuperClot® vs. 4DryField® | -0.6000 | *** | 0.0008 |
| 40 s |  |  |  |
| 4Seal® vs. Starsil® | 0.06667 | ns | >0.9999 |
| 4Seal® vs. PerClot® | 1.600 | **** | <0.0001 |
| 4Seal® vs. ARISTA™ | 1.567 | **** | <0.0001 |
| 4Seal® vs. SuperClot® | 1.900 | **** | <0.0001 |
| 4Seal® vs. 4DryField® | 1.300 | **** | <0.0001 |
| Starsil® vs. PerClot® | 1.533 | **** | <0.0001 |
| Starsil® vs. ARISTA™ | 1.500 | **** | <0.0001 |
| Starsil® vs. SuperClot® | 1.833 | **** | <0.0001 |
| Starsil® vs. 4DryField® | 1.233 | **** | <0.0001 |
| PerClot® vs. ARISTA™ | -0.03333 | ns | >0.9999 |
| PerClot® vs. SuperClot® | 0.3000 | ns | 0.4197 |
| PerClot® vs. 4DryField® | -0.3000 | ns | 0.4197 |
| ARISTA™ vs. SuperClot® | 0.3333 | ns | 0.2306 |
| ARISTA™ vs. 4DryField® | -0.2667 | ns | 0.7379 |
| SuperClot® vs. 4DryField® | -0.6000 | *** | 0.0008 |
| 60 s |  |  |  |
| 4Seal® vs. Starsil® | 0.1333 | ns | >0.9999 |
| 4Seal® vs. PerClot® | 1.600 | **** | <0.0001 |
| 4Seal® vs. ARISTA™ | 1.633 | **** | <0.0001 |
| 4Seal® vs. SuperClot® | 2.000 | **** | <0.0001 |
| 4Seal® vs. 4DryField® | 1.400 | **** | <0.0001 |
| Starsil® vs. PerClot® | 1.467 | **** | <0.0001 |
| Starsil® vs. ARISTA™ | 1.500 | **** | <0.0001 |
| Starsil® vs. SuperClot® | 1.867 | **** | <0.0001 |
| Starsil® vs. 4DryField® | 1.267 | **** | <0.0001 |
| PerClot® vs. ARISTA™ | 0.03333 | ns | >0.9999 |
| PerClot® vs. SuperClot® | 0.4000 | ns | 0.0636 |
| PerClot® vs. 4DryField® | -0.2000 | ns | >0.9999 |
| ARISTA™ vs. SuperClot® | 0..3667 | ns | 0.1228 |
| ARISTA™ vs. 4DryField® | -0.2333 | ns | >0.9999 |
| SuperClot® vs. 4DryField® | -0.6000 | *** | 0.0008 |

| Legend: |  |
| --- | --- |
| symbol: | p-value: |
| ns | 0.1234 |
| * | 0.0332 |
| ** | 0.0021 |
| *** | 0.0002 |
| **** | <0.0001 |

Figure S2. Result of two-way ANOVA with Bonferroni's multiple comparisons test for Water Absorption Rate (WAR).

Additional information for Adhesion potential

Table S6. Results of adhesion of hemostatic powders

| Specimen | Force [N] |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- |
|  | 4SEAL® | Starsil® | PerClot® | ARISTA™ | SuperClot® | 4DryField® |
| 1 | 0.39 | 0.30 | 0.09 | 0.05 | 0.05 | 0.10 |
| 2 | 0.35 | 0.29 | 0.11 | 0.05 | 0.05 | 0.09 |
| 3 | 0.32 | 0.25 | 0.09 | 0.05 | 0.05 | 0.09 |
| Mean | 0.353 | 0.280 | 0.097 | 0.05 | 0.05 | 0.093 |
| Standard Deviation | 0.035 | 0.027 | 0.012 | 0.000 | 0.000 | 0.006 |

Table S7. Result of one-way ANOVA with Bonferroni's multiple comparisons test for Adhesion potential

| Bonferroni's multiple comparisons test | Mean Diff. | Summary | Adjusted P Value |
| --- | --- | --- | --- |
| 4Seal® vs. Starsil® | 0.07333 | ** | 0.0065 |
| 4Seal® vs. PerClot® | 0.2567 | **** | <0.0001 |
| 4Seal® vs. ARISTA™ | 0.3033 | **** | <0.0001 |
| 4Seal® vs. SuperClot® | 0.3033 | **** | <0.0001 |
| 4Seal® vs. 4DryField® | 0.26 | **** | <0.0001 |
| Starsil® vs. PerClot® | 0.1833 | **** | <0.0001 |
| Starsil® vs. ARISTA™ | 0.23 | **** | <0.0001 |
| Starsil® vs. SuperClot® | 0.23 | **** | <0.0001 |
| Starsil® vs. 4DryField® | 0.1867 | **** | <0.0001 |
| PerClot® vs. ARISTA™ | 0.04667 | ns | 0.1499 |
| PerClot® vs. SuperClot® | 0.04667 | ns | 0.1499 |
| PerClot® vs. 4DryField® | 0.003333 | ns | >0.9999 |
| ARISTA™ vs. SuperClot® | 0 | ns | >0.9999 |
| ARISTA™ vs. 4DryField® | -0.04333 | ns | 0.2248 |
| SuperClot® vs. 4DryField® | -0.04333 | ns | 0.2248 |

| Legend: |  |
| --- | --- |
| symbol | p-value: |
| ns | 0.1234 |
| * | 0.0332 |
| ** | 0.0021 |
| *** | 0.0002 |
| **** | <0.0001 |

Figure S3. Result of one-way ANOVA with Bonferroni's multiple comparisons test for Adhesion potential

HEMOSTATIC POWDERS

Table S8. Information about hemostatic powders

| Name | Manufacturer | REF | LOT |
| --- | --- | --- | --- |
| 4SEAL® | Grena Biomed Ltd. | 1203-HP005 | Z/210505 |
| Starsil® | Hemostat Manufacturing GmbH | SS005 | H190001 |
| PerClot® | CryoLife, Inc. | STA0001 | 180303 |
| ARISTA™ | BD, Inc. | SM0002-USA | 6111801 |
| SuperClot® | AMP | SC0003 | 200502 |
| 4DryField® | PlantTec Medical | SK0005-EU | 200101 |

Figure S4. Morphology of evaluated hemostats. A - 4DryField, B - 4Seal, C - Arista, D - PerClot, E - SuperClot, F – Starsil. Images were taken using Revolve Microscope (ECHO A BICO COMPANY, USA) and ECHO Pro ver. 6.4.2 software at 100 x magnification.