|  | AceSight | eSight 3 | IrisVision Live | Jordy |
| --- | --- | --- | --- | --- |
| Price in DKK excluding tax | 34.500 | 59.500 | 29.500 | 19.988 |
| Manufacturer | Zoomax Technology Co. (USA) | eSight corp. (Canada) | IrisVision Global Inc. (USA) | Enhanced vision (USA) |
| Magnification | From 1.1x to 15x | Up to 24x | From 1x to 12x | Up to 30x (in a 2-meter distance) |
| Autofocus | Yes (can be locked in settings) | Yes | Yes | Yes |
| Camera resolution | 8 megapixels | 21.5 megapixels | 12 megapixels | 720p HD (0.9 megapixels) |
| Screen resolution | 1920 x 1080 | 1024 x 768 OLED | 2560 x 1440 OLED | 1280×720 |
| Field of view | 45 degrees horizontal per eye | 37.5 degrees diagonally per eye | 70 degrees horizontally | N/A |
| Contrast | Black/white Blue/white Black/yellow Blue/yellow Green/black All colours can be inverted | Inverted Black/white Blue/white Blue/yellow Yellow/blue | Inverted Black/white Green Yellow | Inverted Black/white Green/black Yellow/black Yellow/blue |
| Overview in zoom | Yes | Yes | No | Yes |
| Navigation in ”frozen picture” | Yes, with head-movement | Yes, with head-movement | Yes, with head-movement | Yes, but only central zoom |
| Adjustable brightness | Yes | Yes | Yes | Yes |
| Remote control | Yes, with a wire | Yes, with a wire | No | Yes, with a wire |
| Battery life | 4 hours | 6 hours | 3.5 hours | 4 and 8 hours (two battery sizes) |
| HDMI connection | No | Yes | No (streaming will soon be an option) | Yes |
| Head-mounted weight | 360 g | 104 g | 498 g | 236 g |
| Additional features | Contour outlining Reduced image | Tablet for instruction | Line reading Bioptic mode Bubble view Reduced image | Anti-glare shield HDMI cable |
| Spectacle correction | Possible to wear glasses with the HMD | Prescription insertion | Possible to wear glasses with the HMD | Prescription insertion |
| Adjustment of pupillary distance | Yes | Yes | Yes (in the software) | No |
| Accessories |  | Extra batteries can be purchased | Wireless charger and remote control can be purchased | CCTV stand (11.200 DKK inc. tax) |

Supplementary Table A1. Overview illustrating differences between the head-mounted displays

DKK = danish kroners

Supplementary Table A2. Assessment of the performance and satisfaction of tasks wearing the head-mounted displays. Six participants evaluated the devices from 1 (worst) to 10 (best). The baseline values are measurements without wearing a head-mounted display. Outdoor distance tasks were not assessed for IrisVision. The number of evaluations of the task is noted in parentheses. Phase 3-6.

|  |  | Baseline |  | AceSight |  | eSight |  | IrisVision |  | Jordy |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Activity | Importance | Performance | Satisfaction | Performance | Satisfaction | Performance | Satisfaction | Performance | Satisfaction | Performance | Satisfaction |
| Indoor distance tasks | 8.0 (9) | 2.1 (9) | 1.8 (9) | 4.7 (6) | 3.3 (6) | 6.0 (4) | 4.5 (4) | 9.3 (3) | 9.7 (3) | 1.8 (4) | 1.3 (4) |
| Reading | 8.1 (7) | 3.6 (7) | 3.0 (7) | 5.0 (5) | 3.8 (5) | 3.3 (3) | 3.3 (3) | 7.0 (1) | 10.0 (1) | 2.4 (5) | 1.8 (5) |
| Outdoor distance tasks | 9.0 (4) | 4.3 (4) | 3.8 (4) | 6.7 (3) | 6.3 (3) | 7. 3(3) | 5.3 (3) |  |  | 1.0 (1) | 1.0 (1) |
| Near tasks | 8.4 (7) | 3.1 (7) | 2.6 (7) | 4.4 (5) | 4.6 (5) | 1.0 (1) | 2.0 (1) | 3.0 (2) | 5.0 (2) | 2.8 (4) | 2.0 (4) |
| Computer-related tasks | 10.0 (2) | 7.5 (2) | 5.0 (2) | 1.0 (1) | 1.0 (1) | 7.0 (1) | 1.0 (1) | 3.0 (1) | 6.0 (1) | 2.0 (1) | 1.0 (1) |
| Mean score | 8.7 | 4.1 | 3.2 | 4.4 | 3.8 | 4.9 | 3.2 | 5.6 | 7.7 | 2.0 | 1.4 |