SUPPORTING INFORMATION

# Supplementary table1  Linearity of HPLC method

|  | Concentration of TEM (mg/mL) | Area under peak |
| --- | --- | --- |
| 1 | 0.001 | 50.4 |
| 2 | 0.005 | 247.1 |
| 3 | 0.01 | 475.2 |
| 4 | 0.05 | 2316.0 |
| 5 | 0.1 | 4640.0 |
| 6 | 0.2 | 9232.8 |
| 7 | 0.25 | 11860.9 |
| 8 | 0.4 | 18232.4 |

HPLC: High Performance Liquid Chromatography; TEM: Temsirolimus

# Supplementary table2  Detection limit of HPLC method

| Detection limit | NO. | Area under peak | Mean (Area) | RSD (Area) | S/N | Mean (S/N) | RSD (S/N) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 0.00003mg/mL | 1 | 1.6 | 1.63 | 3.535% | 5.2 | 5.13 | 1.125% |
|  | 2 | 1.6 |  |  | 5.1 |  |  |
|  | 3 | 1.7 |  |  | 5.1 |  |  |

HPLC: High Performance Liquid Chromatography; RSD: relative standard deviation

# Supplementary table3  Quantification limit of HPLC method

| Quantification limit | NO. | Area under peak | Mean (Area) | RSD (Area) | S/N | Mean (S/N) | RSD (S/N) |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 0.0001mg/mL | 1 | 5.5 | 5.50 | 3.636% | 18.6 | 17.50 | 6.590% |
|  | 2 | 5.3 |  |  | 16.3 |  |  |
|  | 3 | 5.7 |  |  | 17.6 |  |  |

HPLC: High Performance Liquid Chromatography; RSD: relative standard deviation

# Supplementary table4  System Suitability of HPLC method

| NO. | Area under peak | Mean (Area) | RSD (Area) | Time | Mean (Time) | RSD (Time) | Symmetry factor | R | N |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 2333.3 | 2329.37 | 0.122% | 10.887 | 10.794 | 0.454% | 1.01 | 6.41 | 8973 |
| 2 | 2325.2 |  |  | 10.749 |  |  | 1.01 | 6.45 | 8762 |
| 3 | 2331.7 |  |  | 10.791 |  |  | 1.01 | 6.50 | 8812 |
| 4 | 2329.1 |  |  | 10.781 |  |  | 1.01 | 6.46 | 8794 |
| 5 | 2328.9 |  |  | 10.760 |  |  | 1.01 | 6.40 | 8817 |
| 6 | 2328.0 |  |  | 10.793 |  |  | 1.01 | 6.49 | 8747 |

HPLC: High Performance Liquid Chromatography; RSD: relative standard deviation

# Supplementary table5  Precision of HPLC method

| NO. | Area under peak | Mean (Area) | RSD（1） | RSD (Area) | Time | Mean (Time) | RSD（1） | RSD (Time) |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | 2411.4 | 2390.4 | 1.328% | 0.977% | 10.828 | 10.882 | 0.358% | 0.258% |
| 2 | 2428.0 |  |  |  | 10.845 |  |  |  |
| 3 | 2413.9 |  |  |  | 10.885 |  |  |  |
| 4 | 2352.4 |  |  |  | 10.901 |  |  |  |
| 5 | 2357.9 |  |  |  | 10.934 |  |  |  |
| 6 | 2378.8 |  |  |  | 10.897 |  |  |  |
| NO. | Area under peak | Mean (Area) | RSD（2） |  | Time | Mean (Time) | RSD（2） |  |
| 7 | 2368.7 | 2377.2 | 0.379% |  | 10.870 | 10.878 | 0.130% |  |
| 8 | 2387.6 |  |  |  | 10.897 |  |  |  |
| 9 | 2380.9 |  |  |  | 10.881 |  |  |  |
| 10 | 2372.4 |  |  |  | 10.890 |  |  |  |
| 11 | 2367.2 |  |  |  | 10.865 |  |  |  |
| 12 | 2386.6 |  |  |  | 10.862 |  |  |  |

HPLC: High Performance Liquid Chromatography; RSD: relative standard deviation

# Supplementary table 6  Evaluation of HPLC method Ruggedness

| NO. | Area under peak | Mean (Area) | RSD (Area) | Time | Mean (Time) | RSD (Time) |
| --- | --- | --- | --- | --- | --- | --- |
| The original method | 2407.5 | 2388.27 | 5.365% | 11.195 | 11.146 | 8.769% |
| 276nm | 2343.2 |  |  | 11.195 |  |  |
| 280nm | 2325.6 |  |  | 11.195 |  |  |
| 0.9mL/min | 2669.3 |  |  | 12.031 |  |  |
| 1.1mL/min | 2174.6 |  |  | 10.077 |  |  |
| ACN：H2O=68：32 | 2393.5 |  |  | 12.961 |  |  |
| ACN：H2O=72：28 | 2401.7 |  |  | 9.611 |  |  |
| Column temperature：58℃ | 2398.8 |  |  | 11.143 |  |  |
| Column temperature：62℃ | 2380.2 |  |  | 10.904 |  |  |

HPLC: High Performance Liquid Chromatography; RSD: relative standard deviation; ACN: Acetonitrile

# Supplementary table7  EE of pH-responsive TEM-loaded micelles.

|  | Area under peak (before centrifugation） | Area under peak (after centrifugation） | EE |
| --- | --- | --- | --- |
| pH-responsive TEM nanomicelles | 2116.9 | 2016.7 | 95.27% |

EE: Encapsulation Efficiency; TEM:Temsirolimus