# Supplementary documents RS01X validation

# Appendix 1: Analysis clustered measurements

Appendix Table 1: Paired Samples T-test

| Inhalation profiles | Mean | CI 95% (lower ; upper) | P-value |
| --- | --- | --- | --- |
| IPR Duration – smart DPI Duration (milliseconds) | -2 | -18 ; 14 | 0.827 |
| IPR Volume – smart DPI Volume (milliliters) | 19 | 3 ; 36 | 0.024 |
| IPR PIF – smart DPI PIF (liters/minute) | 0.945 | 0.286 ; 1.603 | 0.005 |

Abbreviations: IPR = inhalation profile recorder, smart DPI = smart dry powder inhaler, PIF = Peak inspiratory flow.

Appendix Figure 1. Population Characteristics; Scatterplot for duration, volume, and peak inspiratory flow (PIF).

# Appendix 2: Analyses within groups

Appendix Table 2.1: Paired Samples T-Test -Healthy

| Inhalations | Mean | CI 95% (lower ; upper) | p-value |
| --- | --- | --- | --- |
| First inhalation |  |  |  |
| IPR Duration – smart DPI Duration (milliseconds) | 41 | -7 ; 88 | 0.092 |
| IPR Volume – smart DPI Volume (milliliters) | -5 | -31 ; 20 | 0.683 |
| IPR PIF – smart DPI PIF (liters/minute) | -0.601 | -1.986 ; 0.784 | 0.384 |
| Second inhalation |  |  |  |
| IPR Duration – smart DPI Duration (milliseconds) | -3 | 98 ; 17 | 0.878 |
| IPR Volume – smart DPI Volume (milliliters) | 8 | -36 ; 52 | 0.704 |
| IPR PIF – smart DPI PIF (liters/minute) | -0.271 | -1.976 ;1.435 | 0.749 |
| Third inhalation |  |  |  |
| IPR Duration – smart DPI Duration (milliseconds) | 50 | 6 ; 94 | 0.027 |
| IPR Volume – smart DPI Volume (milliliters) | 75 | 6 ; 145 | 0.035 |
| IPR PIF – smart DPI PIF(liters/minute) | 1.717 | -0.757 ; 4.191 | 0.168 |

Abbreviations: IPR = inhalation profile recorder, smart DPI = smart dry powder inhaler, PIF = Peak inspiratory flow.

Appendix Table 2.2: Paired Samples T-Test -Asthma

| Inhalations | Mean | CI 95% (lower ; upper) | p-value |
| --- | --- | --- | --- |
| First inhalation |  |  |  |
| IPR Duration – smart DPI Duration (milliseconds) | -3 | -47 ; 42 | 0.898 |
| IPR Volume – smart DPI Volume (milliliters) | 2 | -50 ; 53 | 0.951 |
| IPR PIF – smart DPI PIF (liters/minute) | 0.802 | -0.979 ; 2.583 | 0.365 |
| Second inhalation |  |  |  |
| IPR Duration – smart DPI Duration (milliseconds) | -0.035 | -73 ; 2 | 0.064 |
| IPR Volume – smart DPI Volume (milliliters) | 0.063 | -1 ; 127 | 0.052 |
| IPR PIF – smart DPI PIF (liters/minute) | 0.820 | -1.615 ; 3.255 | 0.497 |
| Third inhalation |  |  |  |
| IPR Duration– smart DPI Duration (milliseconds) | 0.015 | -34 ; 63 | 0.541 |
| IPR Volume – smart DPI Volume (milliliters) | 0.066 | 7 ; 124 | 0.029 |
| IPR PIF – smart DPI PIF (liters/minute) | 2.562 | 0.512 ; 4.612 | 0.016 |

Abbreviations: IPR = inhalation profile recorder, smart DPI = smart dry powder inhaler, PIF = Peak inspiratory flow.

Appendix Table 2.3: Paired Samples T-Test -COPD

| Inhalations | Mean | CI 95% (lower ; upper) | p-value |
| --- | --- | --- | --- |
| First inhalation |  |  |  |
| IPR Duration – smart DPI Duration (milliseconds) | -0.037 | -78 ; 4 | 0.073 |
| IPR Volume – smart DPI Volume (milliliters) | -0.024 | -64 ; 15 | 0.219 |
| IPR PIF – smart DPI PIF (liters/minute) | 0.897 | -0.574 ; 2.367 | 0.222 |
| Second inhalation |  |  |  |
| IPR Duration – smart DPI Duration (milliseconds) | -0.018 | -99 ; 63 | 0.652 |
| IPR Volume – smart DPI Volume (milliliters) | -0.002 | -50 ; 46 | 0.944 |
| IPR PIF – smart DPI PIF (liters/minute) | 2.597 | -0.075–5.269 | 0.056 |
| Third inhalation |  |  |  |
| IPR Duration – smart DPI Duration (milliseconds) | -0.041 | -90 ; 8 | 0.099 |
| IPR Volume – smart DPI Volume (milliliters) | -0.016 | -63 ; 31 | 0.482 |
| IPR PIF – smart DPI PIF (liters/minute) | 0.221 | -1.789 –2.230 | 0.824 |

Abbreviations: IPR = inhalation profile recorder, smart DPI = smart dry powder inhaler, PIF = Peak inspiratory flow.

Appendix Figure 2.1: Scatterplots healthy population per inhalation (duration in seconds, volume in liters and Peak Inspiratory Flow in liters/minute).

Appendix Figure 2.2: Scatterplots asthma population per inhalation (duration in seconds, volume in liters and Peak Inspiratory Flow in liters/minute).

Appendix Figure 2.3: Scatterplots COPD population per inhalation (duration in seconds, volume in liters and Peak Inspiratory Flow in liters/minute).

Appendix Table 2.4: Intraclass Correlation Coefficient -Healthy

| Inhalations | Intraclass correlation coefficient | CI 95% (lower-upper) |
| --- | --- | --- |
| First inhalation |  |  |
| IPR Duration– smart DPI Duration (seconds) | 0.99 | 0.97–0.992 |
| IPR Volume – smart DPI Volume (liters) | 1.00 | 0.99–1.00 |
| IPR PIF – smart DPI PIF (liters/minute) | 0.98 | 0.95 –0.99 |
| Second inhalation |  |  |
| IPR Duration – smart DPI Duration (seconds) | 0.99 | 0.97 –0.99 |
| IPR Volume – smart DPI Volume (liters) | 0.98 | 0.96–0.99 |
| IPR PIF – smart DPI PIF (liters/minute) | 0.97 | 0.95 –0.99 |
| Third inhalation |  |  |
| IPR Duration – smart DPI Duration (seconds) | 0.98 | 0.97–0.99 |
| IPR Volume – smart DPI Volume (liters) | 0.94 | 0.88–0.97 |
| IPR PIF – smart DPI PIF (liters/minute) | 0.93 | 0.87–0.96 |

Abbreviations: IPR = inhalation profile recorder, smart DPI = smart dry powder inhaler, PIF = Peak inspiratory flow.

Appendix Table 2.5: Intraclass Correlation Coefficient -Asthma

| Inhalations | Intraclass correlation coefficient | CI 95% (lower; upper) |
| --- | --- | --- |
| First inhalation |  |  |
| IPR Duration – smart DPI Duration (seconds) | 0.99 | 0.98–0.99 |
| IPR Volume– smart DPI Volume (liters) | 0.97 | 0.95–0.99 |
| IPR PIF – smart DPI PIF (liters/minute) | 0.96 | 0.93–0.98 |
| Second inhalation |  |  |
| IPR Duration – smart DPI Duration (seconds) | 0.98 | 0.97–0.99 |
| IPR Volume– smart DPI Volume (liters) | 0.95 | 0.90–0.98 |
| IPR PIF – smart DPI PIF (liters/minute) | 0.96 | 0.92–0.98 |
| Third inhalation |  |  |
| IPR Duration – smart DPI Duration (seconds) | 0.99 | 0.98–1.00 |
| IPR Volume – smart DPI Volume (liters) | 0.96 | 0.91 –0.98 |
| IPR PIF – smart DPI PIF (liters/minute) | 0.96 | 0.91–0.98 |

Abbreviations: IPR = inhalation profile recorder, smart DPI = smart dry powder inhaler, PIF = Peak inspiratory flow.

Appendix Table 2.6: Intraclass Correlation Coefficient -COPD

| Inhalations | Intraclass correlation coefficient | CI 95% (lower; upper) |
| --- | --- | --- |
| First inhalation |  |  |
| IPR Duration – smart DPI Duration (seconds) | 0.99 | 0.98 –1.00 |
| IPR Volume – smart DPI Volume (liters) | 0.98 | 0.97–0.99 |
| IPR PIF – smart DPI PIF (liters/minute) | 0.97 | 0.93 –0.98 |
| Second inhalation |  |  |
| IPR Duration – smart DPI Duration (seconds) | 0.96 | 0.92–0.98 |
| IPR Volume – smart DPI Volume (liters) | 0.98 | 0.96–0.99 |
| IPR PIF – smart DPI PIF (liters/minute) | 0.93 | 0.85–0.97 |
| Third inhalation |  |  |
| IPR Duration– smart DPI Duration (seconds) | 0.99 | 0.98–1.00 |
| IPR Volume– smart DPI Volume (liters) | 0.97 | 0.93 –0.99 |
| IPR PIF – smart DPI PIF (liters/minute) | 0.95 | 0.89 –0.98 |

Abbreviations: IPR = inhalation profile recorder, smart DPI = smart dry powder inhaler, PIF = Peak inspiratory flow.